## Electrical+Optical
C-frame | T1-Q02-UX |  |  | Optical fibre (attenuation) |  |  |  |  | HV | HV | Low voltage
[9 empty] | Cabel con. | 1800V | 7V/4,1 A | 7V/2,1 A
N | Layer | Quarter | Module | Colour | dBm | dB | dBm | dB |  | Imax , nA | "+6V" | "-6V"
1 | L0(X) | 2 (U) | M153 | blue | -32.33 | -5.9 |  |  | OK | 0.3 | 6.28 | -6.65
2 | L0(X) | 2 (U) | M011 | orange | -33.56 | -7.16 |  |  | OK | 0.7 | 6.29 | -6.65
3 | L0(X) | 2 (U) | M006 | green | -30.54 | -4.14 |  |  | OK | 0.4 | 6.26 | -6.65
4 | L0(X) | 2 (U) | M008 | brown | -29.55 | -3.15 |  |  | OK | 0.5 | 6.28 | -6.66
5 | L0(X) | 2 (U) | M009 | grey | -28.18 | -1.76 |  |  | OK | 0.3 | 6.29 | -6.65
6 | L0(X) | 2 (U) | M010 | white | -28.59 | -2.18 |  |  | OK | 0.6 | 6.28 | -6.66
7 | L0(X) | 2 (U) | M154 | red | -27.46 | -1.06 |  |  | OK | 0.2 | 6.28 | -6.65
8 | L0(X) | 2 (U) | WAR004 | blac2 | -26.78 | -0.38 |  |  | OK | 0.6 | 6.28 | -6.65
9 | L0(X) | 2 (U) | WAR089 | yellow | -29.05 | -2.65 |  |  | OK | 0.3 | 6.3 | -6.65
|  |  |  | max | -26.78 | -0.38 | 0 | 0 |  | 0.2 | 6.3 | -6.65
|  |  |  | min | -33.56 | -7.16 | 0 | 0 |  | 0.7 | 6.26 | -6.66
-6H<-->-6R
1 | L0(X) | 0 (L) | M153 | blue | -25.15 | 1.26 |  |  | OK | 0.4 | 6.18 | -6.64
2 | L0(X) | 0 (L) | M011 | orange | -25.89 | 0.52 |  |  | OK | 0.5 | 6.26 | -6.64
3 | L0(X) | 0 (L) | M006 | green | -23.82 | 2.58 |  |  | OK | 0.4 | 6.29 | -6.63
4 | L0(X) | 0 (L) | M008 | brown | -26.18 | 0.22 |  |  | OK | 0.4 | 6.29 | -6.65
5 | L0(X) | 0 (L) | M009 | grey | -25.62 | 0.78 |  |  | OK | 0.3 | 6.29 | -6.64
6 | L0(X) | 0 (L) | M010 | white | -24.84 | 1.56 |  |  | OK | 0.4 | 6.25 | -6.64
7 | L0(X) | 0 (L) | M154 | red | -24.86 | 1.55 |  |  | OK | 0.2 | 6.24 | -6.64
8 | L0(X) | 0 (L) | WAR045 | blac2 | -25.28 | 1.13 |  |  | OK | 0.4 | 6.28 | -6.64
9 | L0(X) | 0 (L) | WAR105 | yellow | -25.08 | 1.33 |  |  | OK | 0.4 | 6.29 | -6.64
|  |  |  | max | -26.18 | 0.22 | 0 | 0 |  | 0.2 | 6.18 | -6.65
|  |  |  | min | -23.82 | 2.58 | 0 | 0 |  | 0.5 | 6.29 | -6.63
1 | L1(U) | 2 (U) | M004 | blue | -25.79 | 0.64 |  |  | OK | 0.2 | 6.27 | -6.65
2 | L1(U) | 2 (U) | M155 | orange | -26.32 | 0.09 |  |  | OK | 0.5 | 6.27 | -6.65
3 | L1(U) | 2 (U) | M012 | green | -24.83 | 1.57 |  |  | OK | 0.2 | 6.27 | -6.64
4 | L1(U) | 2 (U) | M013 | brown | -25.77 | 0.67 |  |  | FE-4-P1 | 0.8 | 6.27 | -6.65
5 | L1(U) | 2 (U) | M014 | grey | -23.77 | 2.64 |  |  | OK | 0.2 | 6.24 | -6.63
6 | L1(U) | 2 (U) | M159 | white | -25.92 | 0.49 |  |  | OK | 0.3 | 6.28 | -6.65
7 | L1(U) | 2 (U) | M016 | red | -24.6 | 1.81 |  |  | OK | 0.2 | 6.26 | -6.65
8 | L1(U) | 2 (U) | WAR005 | blac2 | -25.18 | 1.22 |  |  | OK | 0.3 | 6.27 | -6.65
9 | L1(U) | 2 (U) | WAR091 | yellow | -25.07 | 1.34 |  |  | OK | 0.2 | 6.26 | -6.65
|  |  |  | max | -26.32 | 0.09 | 0 | 0 |  | 0.2 | 6.24 | -6.65
|  |  |  | min | -23.77 | 2.64 | 0 | 0 |  | 0.8 | 6.28 | -6.63
-6H<-->-6R
1 | L1(U) | 0 (L) | M004 | blue | -26.4 | 0.01 |  |  | OK | 0.3 | 6.27 | -6.64
2 | L1(U) | 0 (L) | M155 | orange | -26.62 | -0.22 |  |  | OK | 0.5 | 6.29 | -6.64
3 | L1(U) | 0 (L) | M012 | green | -26.66 | -0.26 |  |  | OK | 0.3 | 6.29 | -6.64
4 | L1(U) | 0 (L) | M013 | brown | -24.59 | 1.85 |  |  | OK | 0.4 | 6.29 | -6.64
5 | L1(U) | 0 (L) | M014 | grey | -25.03 | 1.31 |  |  | OK | 0.2 | 6.25 | -6.64
6 | L1(U) | 0 (L) | M159 | white | -24.02 | 2.39 |  |  | OK | 0.4 | 6.29 | -6.64
7 | L1(U) | 0 (L) | M016 | red | -23.99 | 2.43 |  |  | OK | 0.2 | 6.28 | -6.64
8 | L1(U) | 0 (L) | WAR046 | blac2 | -25.75 | 0.65 |  |  | OK | 0.3 | 6.29 | -6.64
9 | L1(U) | 0 (L) | WAR107 | yellow | -25.13 | 1.27 |  |  | FE-9-P2 | 0.2 | 6.29 | -6.62
|  |  |  | max | -26.66 | -0.26 | 0 | 0 |  | 0.2 | 6.25 | -6.64
|  |  |  | min | -23.99 | 2.43 | 0 | 0 |  | 0.5 | 6.29 | -6.62
Note on D5: Iuri Bagaturia: M004 <----> M011
Note on D17: Iuri Bagaturia: M004 <----> M011
Note on K31: Iuri Bagaturia: no conection.
Note on K48: Iuri Bagaturia: no conection.
Note on L31: Iuri Bagaturia: 1800 V --> 1800 V ???
Note on L48: Iuri Bagaturia: 1800 V --> 1500 V

## GasTightness
[11 empty] | Layer > | L0(X) | [8 empty] | L1(U) | [32 empty] | Layer > | L0(X) | [8 empty] | L1(U)
Layer > | L0(X) | L0(X) | L0(X) | L0(X) | L1(U) | L1(U) | L1(U) | L1(U) |  |  | Nmeas. > | 1 | 7 | 3 | 6 | 5 | 6 | 7 | 9 | 8 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | [24 empty] | Nmeas. > | 1 | 7 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 5.4 | 9
Tube > | F-Low | S-Low | F-Up | S-Up | F-Low | S-Low | F-Up | S-Up |  |  | Time | Pressure , mb | [19 empty] | dP/dt , mb/sec | [21 empty] | Time | Pressure , mb | [19 empty] | dP/dt , mb/sec
Time, sec | P , mb | P , mb | P , mb | P , mb | P , mb | P , mb | P , mb | P , mb |  | min | sec | M153 | M011 | M006 | M008 | M009 | M010 | M154 | war004/045 | war089/105 | M004 | M155 | M012 | M013 | M014 | M159 | M016 | war005/046 | war091/107 |  |  | M153 | M011 | M006 | M008 | M009 | M010 | M154 | war004/045 | war089/105 | M004 | M155 | M012 | M013 | M014 | M159 | M016 | war005/046 | war091/107 |  |  |  |  | sec | M153 | M011 | M006 | M008 | M009 | M010 | M154 | war004/045 | war089/105 | M004 | M155 | M012 | M013 | M014 | M159 | M016 | war005/046 | war091/107 |  |  | M153 | M011 | M006 | M008 | M009 | M010 | M154 | war004/045 | war089/105 | M004 | M155 | M012 | M013 | M014 | M159 | M016 | war005/046 | war091/107
0 | 10.6 | 9.3 | 10.4 | 10.5 | 9.6 | 11.8 | 11.5 | 9.9 | 10.2 |  | 0 | 5.2 | 5.4 | 5.4 | 5.5 | 5.4 | 5.3 | 5.7 | 5.4 | 5.5 | 5.4 | 5.5 | 5.5 | 5.4 | 5.3 | 5.4 | 5.4 | 5.6 | 5.7 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0.01 | 0.01 | 0.01 | 0.005 | 0 | 0.01 | 0.01 | 0 |  |  |  |  | 0 | 5.2 | 5.4 | 5.4 | 5.7 | 5.4 | 5.3 | 5.4 | 5.3 | 5.3 | 5.9 | 5.85 | 5.7 | 5.8 | 5.8 | 5.8 | 5.8 | 5.85 | 5.1 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0.01 | 0 | 0.005 | 0 | 0 | 0.03 | 0 | 0 | 0.005 | 0
10 | 10.6 | 9.3 | 10.4 | 10.5 | 9.5 | 11.8 | 11.5 | 9.9 | 10.15 |  | 10 | 5.2 | 5.4 | 5.4 | 5.5 | 5.4 | 5.3 | 5.7 | 5.3 | 5.5 | 5.4 | 5.4 | 5.4 | 5.3 | 5.25 | 5.4 | 5.3 | 5.5 | 5.7 | 10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0 | 0 | 0.01 | 0.005 | 0.01 | 0 | 0 | 0.01 |  |  |  |  | 10 | 5.2 | 5.4 | 5.4 | 5.7 | 5.4 | 5.3 | 5.4 | 5.2 | 5.2 | 5.9 | 5.8 | 5.7 | 5.8 | 5.5 | 5.8 | 5.8 | 5.8 | 5.1 | 10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0.01 | 0.005 | 0 | 0 | 0.01 | -0.02 | 0 | 0 | 0 | 0.01
20 | 10.5 | 9.3 | 10.3 | 10.5 | 9.4 | 11.8 | 11.5 | 9.9 | 10.1 |  | 20 | 5.2 | 5.4 | 5.4 | 5.5 | 5.4 | 5.3 | 5.7 | 5.2 | 5.5 | 5.4 | 5.4 | 5.4 | 5.2 | 5.2 | 5.3 | 5.3 | 5.5 | 5.6 | 20 |  | 0.005 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0.01 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0 |  |  |  |  | 20 | 5.2 | 5.4 | 5.4 | 5.7 | 5.4 | 5.3 | 5.4 | 5.1 | 5.1 | 5.85 | 5.8 | 5.7 | 5.7 | 5.7 | 5.8 | 5.8 | 5.8 | 5 | 20 |  | 0.005 | 0.005 | 0 | 0 | 0 | 0 | 0.005 | 0.01 | 0.01 | 0 | 0.01 | 0.01 | 0 | 0 | 0.01 | 0.01 | 0.01 | 0
30 | 10.5 | 9.3 | 10.3 | 10.5 | 9.3 | 11.7 | 11.4 | 9.9 | 10 |  | 30 | 5.15 | 5.35 | 5.4 | 5.5 | 5.4 | 5.3 | 5.7 | 5.1 | 5.5 | 5.4 | 5.3 | 5.4 | 5.2 | 5.2 | 5.3 | 5.25 | 5.5 | 5.6 | 30 |  | 0.005 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0.01 | 0 | 0 | 0 | 0.01 | 0.01 | 0.01 | 0.005 | 0.005 | 0.01 | 0 |  |  |  |  | 30 | 5.15 | 5.35 | 5.4 | 5.7 | 5.4 | 5.3 | 5.35 | 5 | 5 | 5.85 | 5.7 | 5.6 | 5.7 | 5.7 | 5.7 | 5.7 | 5.7 | 5 | 30 |  | 0.005 | 0.005 | 0 | 0 | 0 | 0 | 0.005 | 0.01 | 0.01 | 0 | 0.005 | 0 | 0.01 | 0.01 | 0 | 0 | 0 | 0
40 | 10.5 | 9.2 | 10.3 | 10.5 | 9.3 | 11.7 | 11.4 | 9.9 | 10 |  | 40 | 5.1 | 5.3 | 5.4 | 5.5 | 5.4 | 5.3 | 5.7 | 5 | 5.5 | 5.4 | 5.3 | 5.3 | 5.1 | 5.1 | 5.25 | 5.2 | 5.4 | 5.6 | 40 |  | 0 | 0 | 0 | 0.01 | 0 | 0.005 | 0 | 0.01 | 0 | 0 | 0.005 | 0 | 0 | 0 | 0.005 | 0 | 0 | 0.01 |  |  |  |  | 40 | 5.1 | 5.3 | 5.4 | 5.7 | 5.4 | 5.3 | 5.3 | 4.9 | 4.9 | 5.85 | 5.65 | 5.6 | 5.6 | 5.6 | 5.7 | 5.7 | 5.7 | 5 | 40 |  | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.01 | 0.01 | 0 | 0.005 | 0.005 | 0 | 0 | 0.01 | 0.005 | 0.01 | 0.01
50 | 10.5 | 9.2 | 10.3 | 10.5 | 9.2 | 11.7 | 11.4 | 9.9 | 9.9 |  | 50 | 5.1 | 5.3 | 5.4 | 5.4 | 5.4 | 5.25 | 5.7 | 4.9 | 5.5 | 5.4 | 5.25 | 5.3 | 5.1 | 5.1 | 5.2 | 5.2 | 5.4 | 5.5 | 50 |  | 0 | 0 | 0 | 0 | 0 | 0.005 | 0.005 | 0.01 | 0 | 0 | 0.005 | 0.01 | 0.01 | 0.01 | 0 | 0 | 0 | 0 |  |  |  |  | 50 | 5.1 | 5.3 | 5.4 | 5.7 | 5.4 | 5.25 | 5.3 | 4.8 | 4.8 | 5.85 | 5.6 | 5.55 | 5.6 | 5.6 | 5.6 | 5.65 | 5.6 | 4.9 | 50 |  | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.01 | 0.005 | 0 | 0 | 0.005 | 0.01 | 0.01 | 0 | 0.005 | 0 | 0
60 | 10.4 | 9.2 | 10.3 | 10.5 | 9.15 | 11.7 | 11.4 | 9.8 | 9.8 | 1 | 60 | 5.1 | 5.3 | 5.4 | 5.4 | 5.4 | 5.2 | 5.65 | 4.8 | 5.5 | 5.4 | 5.2 | 5.2 | 5 | 5 | 5.2 | 5.2 | 5.4 | 5.5 | 60 |  | 0 | 0 | 0.0016667 | 0 | 0.0033333 | 0 | 0.0016667 | 0.01 | 0.0016667 | 0 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | after resoldering | <==== | before resoldering | ====> | 60 | 5.1 | 5.3 | 5.4 | 5.7 | 5.4 | 5.2 | 5.3 | 4.7 | 4.75 | 5.85 | 5.6 | 5.5 | 5.5 | 5.5 | 5.6 | 5.6 | 5.6 | 4.9 | 60 |  | 0 | 0 | 0.0016667 | 0 | 0.0033333 | 0 | 0 | 0.01 | 0.0083333 | 0.0016667 | 0.0066667 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
90 | 10.4 | 9.1 | 10.3 | 10.4 | 8.9 | 11.6 | 11.4 | 9.8 | 9.6 |  | 90 | 5.1 | 5.3 | 5.35 | 5.4 | 5.3 | 5.2 | 5.6 | 4.5 | 5.45 | 5.4 | 5.1 | 5.15 | 4.9 | 4.9 | 5.05 | 5.05 | 5.3 | 5.4 | 90 |  | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0 | 0.0033333 | 0 | 0.0066667 | 0 | 0 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 90 | 5.1 | 5.3 | 5.35 | 5.7 | 5.3 | 5.2 | 5.3 | 4.4 | 4.5 | 5.8 | 5.4 | 5.4 | 5.35 | 5.4 | 5.5 | 5.5 | 5.5 | 4.8 | 90 |  | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0 | 0.0033333 | 0.0033333 | 0.0083333 | 0.0066667 | 0 | 0.0033333 | 0.0033333 | 0.005 | 0.0066667 | 0.0066667 | 0.0033333 | 0.0033333 | 0.0033333
120 | 10.3 | 9 | 10.2 | 10.4 | 8.8 | 11.5 | 11.3 | 9.75 | 9.4 | 2 | 120 | 5 | 5.2 | 5.3 | 5.3 | 5.3 | 5.1 | 5.6 | 4.3 | 5.45 | 5.4 | 4.95 | 5 | 4.8 | 4.8 | 4.95 | 4.95 | 5.2 | 5.3 | 120 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.0033333 | 0.0066667 | 0.0016667 | 0 | 0.0033333 | 0.0033333 | 0.005 | 0.005 | 0.005 | 0.0016667 | 0.0033333 | 0.0033333 |  |  |  |  | 120 | 5 | 5.2 | 5.3 | 5.6 | 5.3 | 5.1 | 5.2 | 4.15 | 4.3 | 5.8 | 5.3 | 5.3 | 5.2 | 5.2 | 5.3 | 5.4 | 5.4 | 4.7 | 120 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0083333 | 0.0083333 | 0 | 0.0033333 | 0.005 | 0.0066667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
150 | 10.2 | 8.9 | 10.2 | 10.3 | 8.5 | 11.3 | 11.3 | 9.7 | 9.2 |  | 150 | 5 | 5.2 | 5.3 | 5.3 | 5.3 | 5.1 | 5.5 | 4.1 | 5.4 | 5.4 | 4.85 | 4.9 | 4.65 | 4.65 | 4.8 | 4.9 | 5.1 | 5.2 | 150 |  | 0.0033333 | 0.0033333 | 0 | 0 | 0 | 0 | 0 | 0.0066667 | 0 | 0 | 0.005 | 0.0033333 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 150 | 5 | 5.2 | 5.3 | 5.6 | 5.3 | 5.1 | 5.2 | 3.9 | 4.05 | 5.8 | 5.2 | 5.15 | 5 | 5.1 | 5.2 | 5.3 | 5.3 | 4.6 | 150 |  | 0.0033333 | 0.0033333 | 0 | 0 | 0 | 0 | 0.0033333 | 0.0066667 | 0.0083333 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
180 | 10.1 | 8.9 | 10.2 | 10.3 | 8.4 | 11.2 | 11.3 | 9.7 | 9.05 | 3 | 180 | 4.9 | 5.1 | 5.3 | 5.3 | 5.3 | 5.1 | 5.5 | 3.9 | 5.4 | 5.4 | 4.7 | 4.8 | 4.5 | 4.5 | 4.7 | 4.8 | 5 | 5.1 | 180 |  | 0 | 0 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0066667 | 0 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 180 | 4.9 | 5.1 | 5.3 | 5.6 | 5.3 | 5.1 | 5.1 | 3.7 | 3.8 | 5.8 | 5.1 | 5.05 | 4.9 | 4.95 | 5.1 | 5.2 | 5.2 | 4.5 | 180 |  | 0 | 0 | 0.0033333 | 0 | 0.0016667 | 0.0033333 | 0 | 0.0066667 | 0.0066667 | 0 | 0.0033333 | 0.0033333 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
210 | 10.1 | 8.8 | 10.1 | 10.3 | 8.2 | 11.1 | 11.2 | 9.6 | 8.9 |  | 210 | 4.9 | 5.1 | 5.2 | 5.2 | 5.25 | 5 | 5.4 | 3.7 | 5.4 | 5.35 | 4.65 | 4.7 | 4.4 | 4.4 | 4.6 | 4.7 | 4.9 | 5 | 210 |  | 0.0033333 | 0.0016667 | 0 | 0 | 0.0016667 | 0 | 0 | 0.0066667 | 0 | 0 | 0.005 | 0.0016667 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 210 | 4.9 | 5.1 | 5.2 | 5.6 | 5.25 | 5 | 5.1 | 3.5 | 3.6 | 5.8 | 5 | 4.95 | 4.75 | 4.8 | 5 | 5.1 | 5.1 | 4.4 | 210 |  | 0.0033333 | 0.0016667 | 0 | 0.0016667 | 0.0016667 | 0 | 0.0016667 | 0.0066667 | 0.0066667 | 0 | 0.0033333 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
240 | 10 | 8.7 | 10.1 | 10.2 | 8.1 | 11 | 11.2 | 9.6 | 8.7 | 4 | 240 | 4.8 | 5.05 | 5.2 | 5.2 | 5.2 | 5 | 5.4 | 3.5 | 5.4 | 5.35 | 4.5 | 4.65 | 4.25 | 4.3 | 4.5 | 4.6 | 4.8 | 4.9 | 240 |  | 0 | 0.0016667 | 0 | 0.0033333 | 0 | 0 | 0.0016667 | 0.01 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0016667 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 240 | 4.8 | 5.05 | 5.2 | 5.55 | 5.2 | 5 | 5.05 | 3.3 | 3.4 | 5.8 | 4.9 | 4.8 | 4.6 | 4.7 | 4.9 | 5 | 5 | 4.3 | 240 |  | 0 | 0.0016667 | 0 | 0.0016667 | 0 | 0 | 0.0016667 | 0.0066667 | 0.0066667 | 0 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[11 empty] | 270 | 4.8 | 5 | 5.2 | 5.1 | 5.2 | 5 | 5.35 | 3.2 | 5.35 | 5.3 | 4.4 | 4.6 | 4.1 | 4.2 | 4.4 | 4.5 | 4.7 | 4.8 | 270 |  | 0 | 0 | 0.0016667 | 0 | 0 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0016667 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0016667 | 0 |  |  |  |  | 270 | 4.8 | 5 | 5.2 | 5.5 | 5.2 | 5 | 5 | 3.1 | 3.2 | 5.8 | 4.75 | 4.7 | 4.5 | 4.6 | 4.8 | 4.9 | 4.9 | 4.2 | 270 |  | 0 | 0 | 0.0016667 | 0 | 0 | 0.0016667 | 0.0016667 | 0.0066667 | 0.0033333 | 0 | 0.005 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[10 empty] | 5 | 300 | 4.8 | 5 | 5.15 | 5.1 | 5.2 | 4.95 | 5.3 | 3.1 | 5.3 | 5.3 | 4.3 | 4.5 | 4 | 4.1 | 4.35 | 4.4 | 4.65 | 4.8 | 300 |  | 0.0033333 | 0.0033333 | 0.0016667 | 0.0016667 | 0 | 0.0016667 | 0 | 0.0066667 | 0 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.005 | 0.0033333 | 0.0016667 | 0.0033333 |  |  |  |  | 300 | 4.8 | 5 | 5.15 | 5.5 | 5.2 | 4.95 | 4.95 | 2.9 | 3.1 | 5.8 | 4.6 | 4.6 | 4.35 | 4.5 | 4.7 | 4.8 | 4.8 | 4.1 | 300 |  | 0.0033333 | 0.0033333 | 0.0016667 | 0 | 0 | 0.0016667 | 0.0016667 | 0.0066667 | 0.0066667 | 0 | 0.0033333 | 0.0033333 | 0.005 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[11 empty] | 330 | 4.7 | 4.9 | 5.1 | 5.05 | 5.2 | 4.9 | 5.3 | 2.9 | 5.3 | 5.3 | 4.2 | 4.4 | 3.9 | 4 | 4.2 | 4.3 | 4.6 | 4.7 | 330 |  | 0 | 0 | 0 | 0.0016667 | 0 | 0 | 0.0033333 | 0.0033333 | 0 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0 | 0.0016667 | 0.0033333 | 0.0033333 |  |  |  |  | 330 | 4.7 | 4.9 | 5.1 | 5.5 | 5.2 | 4.9 | 4.9 | 2.7 | 2.9 | 5.8 | 4.5 | 4.5 | 4.2 | 4.35 | 4.6 | 4.7 | 4.7 | 4 | 330 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0.005 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[10 empty] | 6 | 360 | 4.7 | 4.9 | 5.1 | 5 | 5.2 | 4.9 | 5.2 | 2.8 | 5.3 | 5.3 | 4.1 | 4.3 | 3.8 | 3.9 | 4.2 | 4.25 | 4.5 | 4.6 | 360 |  | 0.0016667 | 0.0016667 | 0 | 0 | 0 | 0.0033333 | 0 | 0.0066667 | 0 | 0 | 0.0033333 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 |  |  |  |  | 360 | 4.7 | 4.9 | 5.1 | 5.5 | 5.2 | 4.9 | 4.9 | 2.55 | 2.75 | 5.8 | 4.4 | 4.4 | 4.1 | 4.2 | 4.5 | 4.6 | 4.6 | 3.9 | 360 |  | 0.0016667 | 0.0016667 | 0 | 0.0016667 | 0 | 0.0033333 | 0.0016667 | 0.005 | 0.005 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[11 empty] | 390 | 4.65 | 4.85 | 5.1 | 5 | 5.2 | 4.8 | 5.2 | 2.6 | 5.3 | 5.3 | 4 | 4.2 | 3.65 | 3.8 | 4.1 | 4.2 | 4.4 | 4.5 | 390 |  | 0.0016667 | 0.0016667 | 0 | 0.0033333 | 0.0016667 | 0 | 0.0033333 | 0.0033333 | 0.0016667 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 |  |  |  |  | 390 | 4.65 | 4.85 | 5.1 | 5.45 | 5.2 | 4.8 | 4.85 | 2.4 | 2.6 | 5.8 | 4.3 | 4.3 | 4 | 4.1 | 4.4 | 4.5 | 4.5 | 3.8 | 390 |  | 0.0016667 | 0.0016667 | 0 | 0.0016667 | 0.0016667 | 0 | 0.0016667 | 0.0066667 | 0.0033333 | 0 | 0.0033333 | 0.0033333 | 0.0066667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[10 empty] | 7 | 420 | 4.6 | 4.8 | 5.1 | 4.9 | 5.15 | 4.8 | 5.1 | 2.5 | 5.25 | 5.3 | 3.9 | 4.1 | 3.55 | 3.7 | 4 | 4.1 | 4.3 | 4.4 | 420 |  | 0 | 0 | 0.0033333 | 0 | 0.0016667 | 0.0033333 | 0 | 0.0033333 | 0 | 0 | 0.0016667 | 0.0033333 | 0.005 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 |  |  |  |  | 420 | 4.6 | 4.8 | 5.1 | 5.4 | 5.15 | 4.8 | 4.8 | 2.2 | 2.5 | 5.8 | 4.2 | 4.2 | 3.8 | 4 | 4.3 | 4.4 | 4.4 | 3.7 | 420 |  | 0 | 0 | 0.0033333 | 0 | 0.0016667 | 0.0033333 | 0 | 0.0033333 | 0.0066667 | 0 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0016667 | 0
[11 empty] | 450 | 4.6 | 4.8 | 5 | 4.9 | 5.1 | 4.7 | 5.1 | 2.4 | 5.25 | 5.3 | 3.85 | 4 | 3.4 | 3.6 | 3.9 | 4 | 4.25 | 4.3 | 450 |  | 0.0033333 | 0.0033333 | 0 | 0.0016667 | 0 | 0.0016667 | 0 | 0.0066667 | 0.0016667 | 0 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 |  |  |  |  | 450 | 4.6 | 4.8 | 5 | 5.4 | 5.1 | 4.7 | 4.8 | 2.1 | 2.3 | 5.8 | 4.1 | 4.15 | 3.7 | 3.9 | 4.2 | 4.35 | 4.35 | 3.7 | 450 |  | 0.0033333 | 0.0033333 | 0 | 0 | 0 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0033333 | 0 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0016667 | 0.0033333
[10 empty] | 8 | 480 | 4.5 | 4.7 | 5 | 4.85 | 5.1 | 4.65 | 5.1 | 2.2 | 5.2 | 5.3 | 3.8 | 3.95 | 3.3 | 3.5 | 3.8 | 3.9 | 4.2 | 4.2 | 480 |  | 0 | 0 | 0 | 0.0016667 | 0 | 0.0016667 | 0.0033333 | 0.0033333 | 0 | 0 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0016667 |  |  |  |  | 480 | 4.5 | 4.7 | 5 | 5.4 | 5.1 | 4.65 | 4.75 | 2 | 2.2 | 5.8 | 4 | 4.1 | 3.6 | 3.8 | 4.1 | 4.3 | 4.3 | 3.6 | 480 |  | 0 | 0 | 0 | 0 | 0 | 0.0016667 | 0.0016667 | 0.0033333 | 0.0033333 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[11 empty] | 510 | 4.5 | 4.7 | 5 | 4.8 | 5.1 | 4.6 | 5 | 2.1 | 5.2 | 5.3 | 3.7 | 3.9 | 3.2 | 3.4 | 3.7 | 3.85 | 4.1 | 4.15 | 510 |  | 0.0016667 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0033333 | 0.0016667 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0016667 | 0.0033333 | 0.0016667 |  |  |  |  | 510 | 4.5 | 4.7 | 5 | 5.4 | 5.1 | 4.6 | 4.7 | 1.9 | 2.1 | 5.8 | 3.9 | 4 | 3.5 | 3.7 | 4 | 4.2 | 4.2 | 3.5 | 510 |  | 0.0016667 | 0 | 0 | 0.0016667 | 0 | 0 | 0 | 0.005 | 0.0033333 | 0 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
[10 empty] | 9 | 540 | 4.45 | 4.7 | 5 | 4.8 | 5.1 | 4.6 | 5 | 2 | 5.15 | 5.3 | 3.6 | 3.8 | 3.1 | 3.3 | 3.6 | 3.8 | 4 | 4.1 | 540 | [23 empty] | 540 | 4.45 | 4.7 | 5 | 5.35 | 5.1 | 4.6 | 4.7 | 1.75 | 2 | 5.8 | 3.8 | 3.9 | 3.4 | 3.6 | 3.9 | 4.1 | 4.1 | 3.4 | 540
[21 empty] | 0.01 | 0.35 | 0.15 | 0.41 | 0.08 | 0.08 | 0.04 | 0.14 | 0.18 |  | mb/min | 0.086364 | 0.081818 | 0.036364 | 0.081818 | 0.027273 | 0.081818 | 0.072727 | 0.41818 | 0.031818 | 0.0090909 | 0.22727 | 0.20909 | 0.28182 | 0.23636 | 0.2 | 0.18182 | 0.18182 | 0.18182 | [14 empty] | 0 | 0.44 | 0.1 | 0.42 | 0.14 | 0.1 | 0.06 | 0.18 | 0.18 |  | mb/min | 0.086364 | 0.081818 | 0.036364 | 0.031818 | 0.027273 | 0.081818 | 0.081818 | 0.43182 | 0.4 | 0.018182 | 0.23182 | 0.2 | 0.27273 | 0.25455 | 0.20909 | 0.19091 | 0.20455 | 0.19091
[15 empty] | 0.7 | [15 empty] | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 1.5 | 0.5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 1 | 1.5 | [25 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 1 | 1.5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 1 | 1.5
[21 empty] | sec | M155 2nd | M012 2nd | M013 alone |  |  |  | M016 2nd |  |  | 2: 7 | [14 empty] | sec | M155 2nd | M012 2nd | M013 alone |  |  |  | M016 2nd
[21 empty] | 0 | 5.8 | 5.8 |  |  |  |  | 5.8 |  |  |  | 0 | 14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [14 empty] | 0 | 5.8 | 5.8 |  |  |  |  | 5.8 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[21 empty] | 10 | 5.7 | 5.75 |  |  |  |  | 5.8 |  |  |  | 0.086364 | 0.16364 | 0.10909 | 0.32727 | 0.13636 | 0.49091 | 0.50909 | 0.62727 | 0.015909 | 0.0090909 | 0.45455 | 0.62727 | 1.1273 | 1.1818 | 1.2 | 1.2727 | 0.18182 | 0.27273 | [14 empty] | 10 | 5.7 | 5.75 |  |  |  |  | 5.8 |  |  |  | 0.086364 | 0.16364 | 0.10909 | 0.12727 | 0.13636 | 0.49091 | 0.57273 | 0.43182 | 0.6 | 0.018182 | 0.46364 | 0.6 | 1.0909 | 1.2727 | 1.2545 | 1.3364 | 0.20455 | 0.28636
[19 empty] | 5.5 |  | 20 | 5.7 | 5.7 |  |  |  |  | 5.7 | [35 empty] | 20 | 5.7 | 5.7 |  |  |  |  | 5.7
[19 empty] | 5.2 |  | 30 | 5.65 | 5.7 |  |  |  |  | 5.7 |  |  | dP/dt_prod | 0.14 | 0 | 0.14 | 0 | 0 | 0.32 | 0.11 | 0.6 | 0 | 0.02 | 0.23 | 0.19 | 0.38 | 0.17 | 0.32 | 0.05 | 0.38 | 0.24 | [14 empty] | 30 | 5.65 | 5.7 |  |  |  |  | 5.7 |  |  | dP/dt_prod | 0.14 | 0 | 0.14 | 0 | 0 | 0.32 | 0.11 | 0.6 | 0 | 0.02 | 0.23 | 0.19 | 0.38 | 0.17 | 0.32 | 0.05 | 0.38 | 0.24
[19 empty] | 5 |  | 40 | 5.6 | 5.6 |  |  |  |  | 5.7 | [35 empty] | 40 | 5.6 | 5.6 |  |  |  |  | 5.7
[19 empty] | 4.7 |  | 50 | 5.55 | 5.6 |  |  |  |  | 5.6 |  |  | dP/dt_V | 0.14 | 0 | 0.14 | 0 | 0 | 0.32 | 0.11 | 0.6 | 0 | 0.02 | 0.23 | 0.19 | 0.38 | 0.17 | 0.32 | 0.05 | 0.38 | 0.24 | [14 empty] | 50 | 5.55 | 5.6 |  |  |  |  | 5.6 |  |  | dP/dt_V | 0.14 | 0 | 0.14 | 0 | 0 | 0.32 | 0.11 | 0.6 | 0 | 0.02 | 0.23 | 0.19 | 0.38 | 0.17 | 0.32 | 0.05 | 0.38 | 0.24
[19 empty] | 4.5 |  | 60 | 5.5 | 5.5 | 5.2 |  |  |  | 5.6 |  |  | dP/dt_sum | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.46 | 0.57 | 0.6 | 0 | 0.02 | 0.25 | 0.63 | 0.63 | 0.8 | 1.12 | 1.17 | 0.38 | 0.5 | [14 empty] | 60 | 5.5 | 5.5 | 5.2 |  |  |  | 5.6 |  |  | dP/dt_sum | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.46 | 0.57 | 0.6 | 0.6 | 0.02 | 0.25 | 0.63 | 0.63 | 0.8 | 1.12 | 1.17 | 0.38 | 0.5
[19 empty] | 4.2 |  | 90 | 5.4 | 5.4 | 4.9 |  |  |  | 5.5 |  |  | dP/dt_sum/V | 0.14 | 0.07 | 0.046667 | 0.035 | 0.028 | 0.076667 | 0.081429 | 0.4 | 0 | 0.02 | 0.125 | 0.21 | 0.1575 | 0.16 | 0.18667 | 0.16714 | 0.38 | 0.33333 | [14 empty] | 90 | 5.4 | 5.4 | 4.9 |  |  |  | 5.5 |  |  | dP/dt_sum/V | 0.14 | 0.07 | 0.046667 | 0.035 | 0.028 | 0.076667 | 0.081429 | 0.6 | 0.4 | 0.02 | 0.125 | 0.21 | 0.1575 | 0.16 | 0.18667 | 0.16714 | 0.38 | 0.33333
[19 empty] | 4 |  | 120 | 5.25 | 5.3 | 4.65 |  |  |  | 5.4 | [35 empty] | 120 | 5.25 | 5.3 | 4.65 |  |  |  | 5.4
[19 empty] | 3.4 |  | 150 | 5.1 | 5.2 | 4.4 |  |  |  | 5.3 |  |  |  | 0.086364 | 0.077273 | -0.054545 | 0.21818 | -0.19091 | 0.35455 | 0.018182 | 0.41818 | 0.61136 | 0.0090909 | 0.44545 | 0.17273 | 0.5 | 0.054545 | 0.018182 | 0.072727 | 0.18182 | 0.18182 | [14 empty] | 150 | 5.1 | 5.2 | 4.4 |  |  |  | 5.3 |  |  |  | 0.086364 | 0.077273 | -0.054545 | 0.018182 | 0.0090909 | 0.35455 | 0.081818 | 0.43182 | 0.33636 | 0.018182 | 0.44545 | 0.13636 | 0.49091 | 0.18182 | -0.018182 | 0.081818 | 0.20455 | 0.16364
[19 empty] | 2.9 |  | 180 | 5 | 5.1 | 4.2 |  |  |  | 5.2 |  |  |  | 0.01 | 0.022361 | 0.036056 | 0.05 | 0.064031 | 0.078102 | 0.092195 | 0.015 | 0.015811 | 0.01 | 0.022361 | 0.036056 | 0.05 | 0.064031 | 0.078102 | 0.092195 | 0.01 | 0.018028 | [14 empty] | 180 | 5 | 5.1 | 4.2 |  |  |  | 5.2 |  |  |  | 0.01 | 0.022361 | 0.036056 | 0.05 | 0.064031 | 0.078102 | 0.092195 | 0.01 | 0.018028 | 0.01 | 0.022361 | 0.036056 | 0.05 | 0.064031 | 0.078102 | 0.092195 | 0.01 | 0.018028
[19 empty] | 2.4 |  | 210 | 4.9 | 5 | 3.95 |  |  |  | 5.05 | [35 empty] | 210 | 4.9 | 5 | 3.95 |  |  |  | 5.05
[19 empty] | 2 |  | 240 | 4.8 | 4.85 | 3.7 |  |  |  | 5 | [35 empty] | 240 | 4.8 | 4.85 | 3.7 |  |  |  | 5
[19 empty] | 1.7 |  | 270 | 4.7 | 4.7 | 3.5 |  |  |  | 4.9 | [35 empty] | 270 | 4.7 | 4.7 | 3.5 |  |  |  | 4.9
[19 empty] | 1.4 |  | 300 | 4.6 | 4.6 | 3.3 |  |  |  | 4.8 | [35 empty] | 300 | 4.6 | 4.6 | 3.3 |  |  |  | 4.8
[19 empty] | 1.1 |  | 330 | 4.4 | 4.5 | 3.1 |  |  |  | 4.7 | [35 empty] | 330 | 4.4 | 4.5 | 3.1 |  |  |  | 4.7
[19 empty] | 0.9 |  | 360 | 4.3 | 4.4 | 2.9 |  |  |  | 4.6 | [35 empty] | 360 | 4.3 | 4.4 | 2.9 |  |  |  | 4.6
[19 empty] | 0.7 |  | 390 | 4.2 | 4.35 | 2.8 |  |  |  | 4.5 | [35 empty] | 390 | 4.2 | 4.35 | 2.8 |  |  |  | 4.5
[21 empty] | 420 | 4.1 | 4.25 | 2.6 |  |  |  | 4.4 | [35 empty] | 420 | 4.1 | 4.25 | 2.6 |  |  |  | 4.4
[21 empty] | 450 | 4 | 4.2 | 2.4 | [39 empty] | 450 | 4 | 4.2 | 2.4
[21 empty] | 480 | 3.9 | 4.1 | 2.3 | [39 empty] | 480 | 3.9 | 4.1 | 2.3
[21 empty] | 510 | 3.85 | 4 | 2.1 | [39 empty] | 510 | 3.85 | 4 | 2.1
[21 empty] | 540 | 3.8 | 3.9 | 2 | [39 empty] | 540 | 3.8 | 3.9 | 2
[24 empty] | 0.34 | [42 empty] | 0.34
[21 empty] | sec | M155 3rd | [41 empty] | sec | M155 3rd
[21 empty] | 0 | 5.75 | [41 empty] | 0 | 5.75
[21 empty] | 10 | 5.7 | [41 empty] | 10 | 5.7
[21 empty] | 20 | 5.7 | [41 empty] | 20 | 5.7
[21 empty] | 30 | 5.6 | [41 empty] | 30 | 5.6
[21 empty] | 40 | 5.6 | [41 empty] | 40 | 5.6
[21 empty] | 50 | 5.5 | [41 empty] | 50 | 5.5
[21 empty] | 60 | 5.5 | [41 empty] | 60 | 5.5
[21 empty] | 90 | 5.3 | [41 empty] | 90 | 5.3
[21 empty] | 120 | 5.2 | [41 empty] | 120 | 5.2
[21 empty] | 150 | 5.1 | [41 empty] | 150 | 5.1
[21 empty] | 180 | 5 | [41 empty] | 180 | 5
[21 empty] | 210 | 4.85 | [41 empty] | 210 | 4.85
[21 empty] | 240 | 4.7 | [41 empty] | 240 | 4.7
[21 empty] | 270 | 4.6 | [41 empty] | 270 | 4.6
[21 empty] | 300 | 4.5 | [41 empty] | 300 | 4.5
[21 empty] | 330 | 4.4 | [41 empty] | 330 | 4.4
[21 empty] | 360 | 4.3 | [41 empty] | 360 | 4.3
[21 empty] | 390 | 4.2 | [41 empty] | 390 | 4.2
[21 empty] | 420 | 4.1 | [41 empty] | 420 | 4.1
[21 empty] | 450 | 4 | [41 empty] | 450 | 4
[21 empty] | 480 | 3.9 | [41 empty] | 480 | 3.9
[21 empty] | 510 | 3.8 | [41 empty] | 510 | 3.8
[21 empty] | 540 | 3.7 | [41 empty] | 540 | 3.7
Note on O4: Iuri Bagaturia: M004 <----> M011
Note on Q4: Iuri Bagaturia: resoldered
Note on T4: Iuri Bagaturia: resoldered
Note on U4: Iuri Bagaturia: After resoldering gastightnes are bad ~0,03mb/sec. The gas pipe-module connections (gluing) are defect;after reparatur OK.
Note on U33: Iuri Bagaturia: 1st measurement WAR004/045, alone.1st indication of problems
Note on X30: Sebastin Bachmann: 2nd measurement taken as a check
Note on X55: Sebastin Bachmann: 3rd measurement taken a day after, since installation was interrupted
Note on Y30: Sebastin Bachmann: 2nd test taken yet another day after (14 sep 2006) because installation was again interrupted
Note on Z30: Sebastin Bachmann: taken a measurement of M013 alone in order to compare DIRECTLY to QA data from production
Note on AD30: Sebastin Bachmann: 2nd measurement taken as a check
Note on AI4: Iuri Bagaturia: M004 <----> M011
Note on BK4: Iuri Bagaturia: M004 <----> M011
Note on BT30: Sebastin Bachmann: 2nd measurement taken as a check
Note on BT55: Sebastin Bachmann: 3rd measurement taken a day after, since installation was interrupted
Note on BU30: Sebastin Bachmann: 2nd test taken yet another day after (14 sep 2006) because installation was again interrupted
Note on BV30: Sebastin Bachmann: taken a measurement of M013 alone in order to compare DIRECTLY to QA data from production
Note on BZ30: Sebastin Bachmann: 2nd measurement taken as a check
Note on CE4: Iuri Bagaturia: M004 <----> M011

## Tabelle3
M004
5.5
5.5
5.45
5.4
5.4
5.4
5.4
5.4
5.35
5.3
5.3
5.3
5.25
5.2
5.2
5.2
5.15
5.1
5.1
5.1
5.1
5
5